| Regional Inspectors’ Checklist for Field Parameters [This checklist is to be completed during regional plant inspections for Field Laboratories, denoted by certification numbers in the 5000s.] |
| --- |
| Facility Name: | Regional Plant Inspector: |
| Permit #: | Regional Inspector Contact #: |
| Field Lab Certification #: | Region: |
| Lab Contact: | Date: |

I.	Check the parameter(s) performed at this site for reporting purposes.

Total Residual Chlorine (TRC)

pH

Temperature (TEMP)

Dissolved Oxygen (DO)

Specific Conductance (SC)

Settleable Residue (SETT)

II.	General Laboratory (note any exceptions in section XI)

| Are instruments, meters, probes, photometric cells, etc. maintained in good condition? | Yes No |
| --- | --- |
| Are standards, reagents and consumables used within manufacturer expiration dates? [TRC gel standard is exempt.] | Yes No |
| Are the following items documented (√ where applicable): |
| Item | TRC | pH | TEMP | DO | SC | SETT |
| Date of sample collection* |  |  |  |  |  |  |
| Time of sample collection* |  |  |  |  |  |  |
| Sample collector’s initials or signature |  |  |  |  |  |  |
| Date of sample analysis* |  |  |  |  |  |  |
| Time of sample analysis* |  |  |  |  | N/A |  |
| Analyst initials or signature |  |  |  |  |  |  |
| Sample site (i.e., facility name, location, ID, etc.) |  |  |  |  |  |  |
| Instrument ID |  |  |  |  |  | N/A |
| Parameter |  |  |  |  |  |  |
| Data qualifiers, when required |  |  |  |  |  |  |

*Date and time of sample collection and analysis may be the same for in situ or on-site measurements.

| III. Total Residual Chlorine – reference method: |  |
| --- | --- |
| Total Residual Chlorine meter make and model: |  |
| Is a check standard analyzed each day of use? (Circle one: gel or liquid standard) | Yes No |
| Does the check standard recover within ±10% of the known value? | Yes No |
| What is the assigned/observed value of the daily check standard? |  |
| Is a 5-standard calibration verification performed? Note date of last verification: | Yes No |
| Alternatively, does the lab construct a linear regression, using 5 standards, to calculate results? Note date of last calibration curve constructed: | Yes No |
| True values: µg/L mg/L _______ _______ _______ _______ _______ |  |
| Obtained values: µg/L mg/L _______ _______ _______ _______ _______ |  |
| What program are samples analyzed on? |  |
| Are results reported in proper units? Check one: µg/L mg/L | Yes No |
| Are results reported between the facility’s permit limit and the compliance limit of 50 µg/L? | Yes No |
| Are results less than the low standard reported as “<x”, where x=low standard conc.? | Yes No |
| Is DPD/buffer added within 15 minutes of collection? | Yes No |
| Is a post-analysis check standard analyzed if samples are analyzed at multiple sites? | Yes No |
| IV. pH – reference method: |  |
| pH meter manufacturer and model: |  |
| Is the pH meter calibrated with at least 2 buffers per mfg’s instructions each day of use? Note buffers used: | Yes No |
| Is the pH meter calibration checked with an additional buffer prior to sample analysis? Note check buffer used: | Yes No |
| Does the check buffer read within ±0.1 S.U. of the known value? | Yes No |
| Are the following items documented: |  |
| Meter calibration? | Yes No |
| Check buffer reading(s)? | Yes No |
| Are samples analyzed within 15 minutes of collection? | Yes No |
| Are sample results reported to 0.1 pH units on the eDMR? | Yes No |
| Is a post-analysis check buffer analyzed if samples are analyzed at multiple sites? | Yes No |

| V. Temperature – reference method: |  |
| --- | --- |
| What instrument(s) is used to measure temperature? Check all that apply: pH meter DO meter Conductivity meter Digital thermometer Glass thermometer |  |
| Is the instrument/thermometer calibration checked at least annually against a NIST traceable thermometer? | Yes No |
| Is NIST traceability documentation maintained on site? | Yes No |
| Are samples measured within 15 minutes? | Yes No |
| Are sample results reported in degrees C? | Yes No |
| VI. Dissolved Oxygen – reference method: |  |
| DO meter make and model: |  |
| Is the air calibration of the DO meter performed each day of use? | Yes No |
| Is meter calibration documented? | Yes No |
| Are samples analyzed within 15 minutes of collection? | Yes No |
| Are results reported in mg/L? | Yes No |
| If samples are analyzed at multiple sites, is the meter recalibrated at each site or a post-analysis calibration verification performed? | Yes No |
| Does the post-analysis verification theoretical value agree within 0.5 mg/L of the meter calibration reading? | Yes No |
| VII. Conductivity – reference method: |  |
| Conductivity meter make and model: |  |
| Is the meter calibrated daily according to the manufacturer’s instructions? Note standard used (this is generally a one-point calibration): | Yes No |
| Is a daily check standard analyzed? Note value: | Yes No |
| Is meter calibration documented? | Yes No |
| Are samples analyzed within 28 days of collection? | Yes No |
| Is a post-analysis check standard analyzed if multiple samples are analyzed? |  |
| Are results reported in µmhos/cm (some meters display equivalent µS/cm units)? | Yes No |
| VIII. Settleable Residue – reference method: |  |
| Does the laboratory have an Imhoff Cone in good condition? | Yes No |
| Is the sample settled for 1 hour? | Yes No |
| Is the sample agitated after 45 minutes? | Yes No |
| Are the following items documented: |  |
| Volume of sample analyzed? Note volume analyzed (must use 1L): | Yes No |
| Date and time of sample analysis (settling start time)? | Yes No |
| Time of agitation after 45 minutes of settling? | Yes No |
| Sample analysis completion (settling end time)? | Yes No |
| Are samples analyzed within 48 hours of collection? | Yes No |
| Are results reported in ml/L? | Yes No |
| IX. Was a paper trail (comparing contract lab and on-site data to DMRs) performed? If so, list months reviewed: | Yes No |
| X. Is follow-up by the Laboratory Certification program recommended? | Yes No |

XI.      Additional comments: ________________________________________________________________________________________

________________________________________________________________________________________

________________________________________________________________________________________

________________________________________________________________________________________

________________________________________________________________________________________

________________________________________________________________________________________

________________________________________________________________________________________

________________________________________________________________________________________

________________________________________________________________________________________

Please submit a copy of this completed form to the Laboratory Certification program at:

DWR Lab Certification, Water Sciences Section, 1623 Mail Service Center, Raleigh NC, 27699-1623

Electronic copies may be emailed to dana.satterwhite@ncdenr.gov.

Revision 4/2018